Item Requirements

○ Article Name: Strain/Voltage Monitoring Equipment

○ Type and Quantity

| Item No. | Item | Q'ty |
| --- | --- | --- |
| 1 | 16ch Strain/Voltage Compact Recorder Includes : - Input cable for strain and voltage - Bridge Box (16ch, 350ohm) - AC Adapter - Control & Analysis software - Battery Unit - Remote Control Unit - Carrying Case | 1 set |

○ Latest Delivery Date: 90 days from the Effective Date of Contract

○ Requirements

- General Feature

  - Compact recorder shall support standalone recording and online recording through a LAN or USB connection.

  - Modular type for easy expansion and up to 128 channels. 16 channel strain/voltage measuring unit shall be included

  - Shock resistance: 490 m/s2 (50 G)

  - All channels synchronous 20 kHz (For 4 channels) Measurement of 1 channel at max. 100 kHz

  - Synchronous sampling of all channels shall be supported with minimum frequency 1kHz

  - Rechargeable battery power unit mounted on the main body shall be included to supply electrical power during minimum 1 hour

- Main Unit

  - Data Saving location : PC & SD Card

  - Interface : USB 2.0, Ethernet

  - Measurement Setting method : Online (PC software), Offline (SD card)

  - Sampling : Max 100kHz

  - Recording Mode : Manual, Trigger, Interval

  - Power : 10~30V DC, 100~240VAC (by AC Adaptor), Battery unit

  - Vibration resistance : 49 m/s2 (5 G)

  - Shock resistance   : 490 m/s2 (50 G)

- Strain/Voltage Unit

  - Measuring targets : Strain gage, Strain gage transducer, Voltage

  - Bridge/Sensor Excitation : 2.5V DC

  - Measuring range

    - Strain ; Minimum 1000 X10-6 strain , Maximum 50000 X10-6 strain

    - Voltage ; Minimum  1V, Maximum 50V

  - Accuracy : within ±0.1% FS

- Battery Unit

  - Battery : Li-ion battery

  - Input DC : 10~30V DC

  - Output   : 10V DC

  - Discharging time : during 2.5A (25w) output ; 60min or more

- Control Software

  - Y-time, X-Y, bar graphs, circular meters, and numeric display are possible

  - Measuring condition setting, data acquisition, data collection, data reproduction and file conversion. For analyzing the acquired data

- Analysis Software

  - Allow a data file to be processed for waveform display, FFT analysis, statistic operations, header information display, tabulation of numeric dataand setting display conditions.

  - Extracts and converts the data files into CSV files.

  - Arithmetic operations, Histogram analysis, Filtering, Differentiation & integration

- Remote Control Unit

  - Operation Button : REC, STOP, BAL, OPTION1, OPTION2

- Carrying Case

  - Inner Size : 55.3 X 42.4 X 27 cm

  - IP67 / Def Stan 81-41/STANAG 4280

- All Standard accessories should be supplied

- One year warranty

○ Payment terms

Payment for the contract shall be made as follows:

| Ratio | Payment terms |
| --- | --- |
| 100% | Receipt of Invoice and Final Acceptance of products by KARI |