Item Requirements

○ Article Name: 4-Port Remote Spectrum Analyzer with Software

○ Type and Quantity

| Item No. | Item | Q'ty |
| --- | --- | --- |
| 1 | LPT-3000RX4 (1 RU, DSP Technology-Based 9 kHz to 3 GHz Remote Control Monitoring Spectrum Analyzer with 4 Switchable RF Input Ports) - SPECTRAS-GEO Yearly Support - SPECTRAS-GEO - Shipping & Handle | 3 |

○ Latest Delivery Date: 50 days from the Effective Date of Contract

○ Requirements

- Frequency

- Frequency

  - Range: 9 kHz to 3 GHz

  - Resolution: 1 Hz

- Frequency Reference

  - Accuracy: +/- (period since last adjustment X aging rage) + stability over temperature + supply voltage stability

  - Aging Rate: +/- 1ppm max.

  - Frequency stability over Temp,: +/- 0.025 ppm (o to 50 °C)

- Frequency Readout Accuracy

  - Start, Stop, Center Marker: +/-(marker frequency indication *frequency reference accuracy +10% * RBW +frequency resolution

  - Trace points: Max 601, min 6

- Marker Frequency Counter

  - Resolution: 1 Hz, 10 Hz, 100 Hz, 1kHz

  - Accuracy: +/- (marker frequency RBW/Span >=0.02; indication* )

- Frequency Span

  - Range: 0 Hz, 100 Hz to 3 GHz

  - Resolution: 1 Hz

- Phase Noise

  - 10 kHz: <-88 dBc/Hz

  - 100 kHz: <-95 dBc/Hz

  - 1 MHz: <-113dBc/Hz

- Resolution Bandwidth Filter

  - Filter Bandwidth

    - 1 Hz to 1 MHz in 1-3-10 sequence (-3dB bandwidth)

    - 200 Hz, 9 kHz, 120 kHz, 1 MHz (-6dB bandwidth)

  - Accuracy

    - +/- 8%, RBW = 1MHz

    - +/- 5%, RBW < 1 MHz

  - Shape Factor

    - <4.5 : 1

- Video Bandwidth (VBW) Filter

  - Filter bandwidth: 1Hz to 1MHz in 1-3-10 sequence

- Amplitude

- Amplitude Range

  - Measurement range

    - 100 kHz to 1MHz: Display Average Noise Level to 18 dB

    - 1 MHz to 10 MHz: Display Average Noise Level to 21 dB

    - 10 MHz to 3 GHz: Display Average Noise Level to 25 dB

- Attenuator

  - Input Attenuator Range: 0 to 50 dB, in 1dB step

- Maximum Safe Input level

  - Average Total Power: <+27 dB

  - DC Voltage +/- 50 V

- 1dB Gain Compression

  - Total Power at First Mixer: >0dBm

  - Total Power at the Preamp: >-22dBm

  - Mixer power level (dBm) = input power (dBm) – attenuation (dB)

- Display Average Noise Level

  - Preamp off: 0 dB attenuation, RF Input is terminated with a 50 Ohm load. RBW 10 Hz, VBW 10 Hz, span 500 Hz, reference level = -50 dBm, trace average>40

  - 9 KHz to 100 KHz < -88 dBm Nominal

  - 100 KHz to 1 MHz <-85 dBm 3*(f/100kHz) dBm Nominal

  - 1 MHz to 10 MHz < -117 dBm Nominal

  - 10 MHz to 2 GHz <-117 dBm Nominal

  - 2 GHz to 3 GHz <-117 dBm Nominal, base model

  - 2 GHz to 3 GHz <-111 dBm Nominal, CID option

  - Preamp on 0 dB attenuation, RF Input is terminated with a 50 Ω load. RBW 10 Hz, VBW 10 Hz, span 500 Hz, reference level= -60 dBm, trace average ≥ 40

  - 100 kHz to 1 MHz < -103 dBm – 3 *(f/100 kHz) dBm

  - 1 MHz to 10 MHz < -137 dBm Nominal, base model

  - 1 MHz to 10 MHz <-135 dBm Nominal, CID option

  - 10 MHz to 3 GHz <-137 dBm + 3 *(f/1 GHz) dB Nominal

- Absolute Amplitude Accuracy

  - Absolute Point Center = 160 MHz, RBW 10 kHz, VBW 1 kHz, span 100 kHz, log scale, 1 dB/div, peak detector, 23 ± 1°C, Signal at Reference Level

  - Preamp off ± 0.3 dB Ref level 0 dBm, 10 dB RF attenuation

  - Preamp on ± 0.4 dB Ref level -30 dBm, 0 dB RF attenuation

- Frequency Response

  - Preamp off Attenuation: 10 dB, Ref: 160 MHz, 20 to 30°C

  - 100 kHz to 2.0 GHz ± 0.5 dB Nominal, base model

  - 2.0 GHz to 3.0 GHz ± 0.8 dB Nominal, base model

  - 100 kHz to 2.0 GHz ± 2.0 dB Nominal, CID option

  - 2.0 GHz to 3.0 GHz ± 3.0 dB Nominal, CID option

  - Preamp on Attenuation: 0 dB, Ref: 160 MHz, 20 to 30°C

  - 1 MHz to 2 GHz ± 0.6 dB Nominal, base model

  - 2 GHz to 3 GHz ± 0.8 dB Nominal, base model

  - 1 MHz to 2 GHz ± 2.0 dB Nominal, CID option

  - 2 GHz to 3 GHz ± 3.0 dB  Nominal, CID option

- Spurious Response

  - Second Harmonic Intercept: Preamp off, signal input -30 dBm, 0 dB attenuation

    - + 35 dBm Typical, 10 MHz < fc < 775 MHz

    - + 60 dBm Typical, 10 MHz < fc < 1.5 MHz

  - Third-order Intercept  Preamp off, signal input –30 dBm, 0 dB attenuation

    - >1dBm 300 MHz to 3.0 GHz

  - Input Related Spurious: <-60 dBc Input signal level – 30dBm, Att. Mode= 0dB, 20- 30°C

  - Residual Response (inherent): <-90 dBm Input terminated, 0 dB attenuation, Preamp off