Item Requirements

○ Article Name: Gifford-Mcmahon cryocooler for thermometer calibration and experiments of KSLV-II

○ Type and Quantity

| Item No. | Item | Q'ty |
| --- | --- | --- |
| 1 | Gifford Mcmahon Cryocooler + compressor set Cryomech AL60 | 1 set |

○ Delivery Date : Within July. 31, 2019

○ Requirement:

-Cold head unit+compressor set required.

-Compressor package (air- cooled)

-Cooling capacity in Watts : (60Hz) - 40W @ 50K        49W @ 60K

(50 Hz)- 57W @ 70K        60W @ 80K

-Compressor input power : 208/230 VAC; 1 phase; 60Hz

-Lowest temperature : 25K with no load

-Cool down time : 15 min. to 80K

-Helium Flexible lines : 10Ft (3m) / standard length

-Helium Pressure : ~13.8 bar

-Cold head interface drawing required. (units displayed in mm)

○ Delivery Item

1. Cold head unit

2. Compressor unit

3. Helium line unit.

4. Power connection wire between compressor and cold head

5. Power cable of compressor unit